ELECTRONIC SUPPLEMENTARY MATERIAL

Development of a rapid in-situ analysis method using sheath-flow probe electrospray ionisation-mass spectrometry (sfPESI-MS) for the direct identification of cocaine metabolites in dried blood spots

Ayoung Kima,*, Paul F. Kellya, Matthew A. Turner a and James C. Reynoldsa,*.

a Centre for Analytical Science, Department of Chemistry, Loughborough University, LE11 3TU, UK

* Corresponding author: a.kim2@lboro.as.uk and j.c.reynolds@lboro.ac.uk

- (b)                 (c)

Figure S1. Example of blood spots: (a) 3 sets of deposited dried blood spots (DBS) to analyse, (b) estimation of spatial resolution (0.03 mm2 ≤ area ≤ 0.79 mm2) of the sfPESI probe and (c) actual DBS sampling model using a sfPESI probe

Table S1. Optimised Parameters of Orbitrap MS

| Parameters |  | Values | Units |
| --- | --- | --- | --- |
| Scan parameters | Scan Mass range | 50 - 450 | a m/z |
|  | Mass Resolution | 25,000 | - |
|  | Scan Rate (Average) | 2.4 | b scans/sec |
|  | Polarity | Positive | - |
| Ion source parameters | Capillary Temperature | 300 | °C |
|  | Capillary Voltage | 32.5 | c V |
|  | Tube Lens Voltage | 90.0 | V |
|  | Skimmer Voltage | 28.0 | V |

a m/z: mass-to-charge ratio, b sec: second and c V: volts

Table S2. Comparative data showing the effect of extraction solvent chemical modification using 0.1% formic acid on the replicate (n=10) sfPESI-MS analysis of a 5 μg/mL liquid standard solution containing benzoylecgonine (BZE).

| Molecular Ion Form of adducted BZE | Target Mass | Mass Spectrum Intensity (c/s) | Concentration of Formic Acid |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 0 % | 0.1 % | 0.5 % |
| [M+H]+ | m/z 290 | M (n = 10) | 1.42.E+05 | 4.37.E+06 | 2.78.E+06 |
|  |  | STDEV | 1.12.E+05 | 1.88.E+06 | 1.61.E+06 |
|  |  | % RSD | 79 | 43 | 58 |
| [M+Na]+ | m/z 312 | M (n = 10) | 4.33.E+05 | 6.74.E+05 | 2.35.E+05 |
|  |  | STDEV | 3.11.E+05 | 4.32.E+05 | 1.69.E+05 |
|  |  | % RSD | 72 | 64 | 72 |
| [M+K]+ | m/z 328 | M (n = 10) | 7.79.E+04 | 1.79.E+05 | 5.84.E+04 |
|  |  | STDEV | 6.23.E+04 | 8.48.E+04 | 3.22.E+04 |
|  |  | % RSD | 80 | 47 | 55 |

M: mean or average, n: the number of measurements, STDEV: standard deviation and % RSD: relative standard deviation

Table S3. Comparative data showing the effect of extraction solvent chemical modification using 0.5 mM sodium acetate on the replicate (n=10) sfPESI-MS analysis of a 5 μg/mL liquid standard solution containing benzoylecgonine (BZE).

| Molecular Ion Form of adducted BZE |  | [M+H]+ |  | [M+Na]+ |  | [M+K]+ |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Target Mass |  | m/z 290 |  | m/z 312 |  | m/z 328 |  |  |
| Concentration of Sodium Acetate |  | 0 mM | 0.5 mM | 0 mM | 0.5 mM | 0 mM | 0.5 mM |  |
| Mass Spectrum Intensity (c/s) | M (n = 10) | 1.42.E+05 | 1.38.E+04 | 4.33.E+05 | 1.20.E+06 | 7.79.E+04 | 4.45.E+04 |  |
|  | STDEV | 1.12.E+05 | 3.31.E+03 | 3.11.E+05 | 3.84.E+05 | 6.23.E+04 | 1.69.E+04 |  |
|  | % RSD | 79 | 24 | 72 | 32 | 80 | 38 |  |

M: mean or average, n: the number of measurements, STDEV: standard deviation and % RSD: relative standard deviation

mM(10-3 mol/L): millimolar (Molar concentration is units of moles of solute per litre of solution)

Table S4. Comparative data showing the effect of extraction solvent chemical modification using 0.5 mM sodium acetate and 0.1% formic acid on the replicate (n=10) sfPESI-MS analysis of a 5 μg/mL liquid standard solution containing benzoylecgonine (BZE), ecgonine methyl ester (EME) and cocaethylene (CE).

| Molecular Ion Form |  | [BZE+H]+ |  | [BZE+Na]+ |  | [EME+H]+ |  | [CE+H]+ |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Target Mass |  | m/z 290 |  | m/z 312 |  | m/z 200 |  | m/z 318 |  |  |
| Chemical Modifier |  | None | 0.1% FA 0.5mM SA | None | 0.1% FA 0.5mM SA | None | 0.1% FA 0.5mM SA | None | 0.1% FA 0.5mM SA |  |
| Mass Spectrum Intensity (c/s) | M (n=10) | 6.21E+05 | 2.07E+06 | 1.23E+05 | 4.09E+05 | 1.57E+06 | 3.54E+06 | 8.54E+05 | 1.88E+06 |  |
|  | STDEV | 5.04E+05 | 6.66E+05 | 8.03E+04 | 1.58E+05 | 8.65E+05 | 1.07E+06 | 8.42E+05 | 6.84E+05 |  |
|  | % RSD | 81 | 32 | 65 | 39 | 55 | 30 | 99 | 36 |  |

M: mean or average, n: the number of measurements, STDEV: standard deviation and % RSD: relative standard deviation

mM(10-3 mol/L): millimolar (Molar concentration is units of moles of solute per litre of solution)

Figure S2. An Example of calibration obtained for the [M+H]+ ion of cocaethylene (m/z 318.17) each data point is calculated from 10 replicate measurements (n=10): LoD = 0.15 μg/mL and R2 = 0.9948.

Figure S3. An Example of calibration obtained for the [M+H]+ ion of ecgonine methyl ester (m/z 200.13) each data point is calculated from 10 replicate measurements (n=10): LoD = 1.31 μg/mL and R2 = 0.9895.

*The circled point at 1 μg/mL was determined to be below the experimental limit-of-detection as determined by 3σ/m (σ: the standard deviation in the sample response at the lowest sample concentration and m: the slope value of the linear correlation graph).

Table S5. Calibration and reproducibility data obtained from replicate dried blood spot analyses (n=10) for ecgonine methyl ester (EME), benzoylecgonine (BZE) and cocaethylene (CE)

| Metabolite | Concentration | 0 µg/mL | 0.5 µg/mL | 1 µg/mL | 2 µg/mL | 5 µg/mL | 10 µg/mL |
| --- | --- | --- | --- | --- | --- | --- | --- |
| EME m/z 200.13 | M (n=10) | 0 | 7.33E+03 | 1.26E+04 | 1.55E+04 | 4.46E+04 | 7.05E+04 |
|  | STDEV | 0 | 0.00E+00 | 3.36E+03 | 1.91E+03 | 1.31E+04 | 1.05E+04 |
|  | % RSD | 0 | 0 | 27 | 12 | 29 | 15 |
| BZE m/z 290.14 | M (n=10) | 0 | 7.24E+03 | 1.35E+04 | 3.02E+04 | 6.97E+04 | 1.46E+05 |
|  | STDEV | 0 | 1.36E+03 | 3.81E+03 | 8.40E+03 | 1.34E+04 | 2.67E+04 |
|  | % RSD | 0 | 19 | 28 | 28 | 19 | 18 |
| CE m/z 318.17 | M (n=10) | 0 | 4.03E+04 | 1.00E+05 | 2.05E+05 | 5.50E+05 | 9.78E+05 |
|  | STDEV | 0 | 5.31E+03 | 1.56E+04 | 3.35E+04 | 1.17E+05 | 1.83E+05 |
|  | % RSD | 0 | 13 | 16 | 16 | 21 | 19 |

M: mean or average, n: the number of measurements, STDEV: standard deviation and % RSD: relative standard deviation